## Sk Low Energy Seed_CLRPv2
CLRP_TG43v2 DB
Air-kerma per history, is scored in vacuo, using a scoring voxel located on the transverse axis 10 cm away from the source and of size 2.66×2.66×0.05 cm3 when representing the solid angle subtended
by the primary collimator of the NIST WAFAC detector and then corrected for lateral and tickness of the scoring voxel to give the air-kerma on the central axis at point 10 cm from the source,s mid-point.
Air-kerma strength per history factors, SKhist, are then calculated by multiplying the air kerma per history at 10 cm by d2 =102.
I-125 seeds |  |  | Pd-103 seeds |  |  | Cs-131 seeds
Number | Source name | [SKhist (Gy cm²/hist)] | Number | Source name | [Skhist (Gy cm² /hist)] | Number | Source name | [Skhist (Gy cm² /hist)]
1 | EchoSeed, 6733 (2µm active layer) | (3.96419+/-0.00035)E-14 | 1 | BestPd-103, 2335 | (6.5805+/-0.0006)E-14 | 1 | Proxcelan, CS-1 Rev2 | (3.71551+/-0.00028)E-14
2 | EchoSeed, 6733 (10µm active layer) | (3.72119+/-0.00033)E-14 | 2 | BrachySeed, Pd-1 | (5.7029+/-0.0005)E-14
3 | OncoSeed, 6702 | (5.7060+/-0.0008)E-14 | 3 | IsoSeed, IsoSeed Pd-103 | (6.2628+/-0.0005)E-14
4 | OncoSeed, 6711 | (3.7666+/-0.0007)E-14 | 4 | InterSource, 1031L | (6.1888+/-0.0005)E-14
5 | ThinSeed, 9011 | (4.1373+/-0.0008)E-14 | 5 | OptiSeed, 1032P | (1.16629+/-0.00007)E-13
6 | Braquibac | (3.49892+/-0.00035)E-14 | 6 | Advantage, IAPd-103A | (8.1498+/-0.0006)E-14
7 | 125Implant Seeds, STM1251 | (3.5102+/-0.0002)E-14 | 7 | Prospera, Med3633 | (8.0819+/-0.0006)E-14
8 | Best I-125, 2301 | (4.13516+/-0.00023)E-14 | 8 | PharmaSeed, BT-103-3 | (7.5920+/-0.0006)E-14
9 | BrachySeed, LS-1 | (5.02309+/-0.00026)E-14 | 9 | TheraSeed, 200 (Light seed) | (6.4261+/-0.0006)E-14
10 | IsoSeed I-125, I25.S06 (.S16) | (4.8312+/-0.0005)E-14 | 10 | TheraSeed, 200 (Heavy seed) | (4.6940+/-0.0005)E-14
11 | IsoSeed I-125, 125.S17 | (3.99444+/-0.00024)E-14 | 11 | CS10 CivaString | (1.146080+/-0.000009)E-13
12 | IsoSeed I-125, I25.S17plus | (4.01410+/-0.00024)E-14 | 12 | CS20 CivaString | (1.143455+/-0.000010)E-13
13 | InterSource 125, 1251L | (4.87264+/-0.00035)E-14 | 13 | CS40 CivaString | (1.132473+/-0.000019)E-13
14 | IsoStar, IS-12501 | (4.17368+/-0.00035)E-14 | 14 | CS50 CivaString | (1.124660+/-0.000015)E-13
15 | IPlant, 3500 | (4.83577+/-0.00035)E-14 | 15 | CT10 CivaThin | (1.149118+/-0.000019)E-13
16 | Advantage I-125, IAI-125A | (4.08677+/-0.00024)E-14 | 16 | CT20 CivaThin | (1.146423+/-0.000019)E-13
17 | ProstaSeed, 125SL | (4.35671+/-0.00032)E-14 | 17 | CT40 CivaThin | (1.135431+/-0.000019)E-13
18 | Prospera, Med3631 | (5.74785+/-0.00028)E-14 | 18 | CT50 CivaThin | (1.127557+/-0.000019)E-13
19 | selectSeed I-125, 130.002 | (4.0042+/-0.0008)E-14
20 | PharmaSeed, BT-125-1 | (3.9590+/-0.0004)E-14
21 | PharmaSeed, BT-125-2 | (3.97254+/-0.00031)E-14
22 | I-Seed I-125, AgX100 | (3.94658+/-0.00034)E-14